Instructions

12wellChamber,removable

The 12 well Chamber, removable is a removable silicone chamber for cell culture and

immunoﬂuorescence stainings. It allows the use of standard cultivation, staining and

mountingtechniqueswithcoverslipsealing.   12wellChamber,removablesupportsup-

rightmicroscopyandlong-termstorageofcellculturemicroscopysamplesaftermount-

ingtheglassslidewithacoverslip1.    12wellChamber,removableisnotrecommended

forlivecellmicroscopyoninvertedmicroscopessincecellsgrowona1mmmicroscopy

glassslide.

1Suitable24mm×60mmcoverslipsareprovidedbyibidi(10811).

Material

SeedingCells

The 12 well Chamber, removable is comprised of a self–

adhesive 12 well silicone gasket mounted on a standard

microscopyglassslide.  Thegasketismanufacturedfrom

biocompatiblesiliconematerial.  Althoughbothmaterials

areautoclavableandcompatiblewithalcohols,wedonot

recommendreusingthem.

Trypsinize and count cells as usual. Dilute the cell sus-

pensiontothedesiredconcentration.  Dependingonyour

celltype,applicationofa4-9×104     cells/mlsuspension

shouldresultinaconﬂuentlayerwithin2-3days.

• Apply250µlcellsuspensionintoeachwell.    Avoid

shakingasthiswillresultininhomogeneousdistri-

butionofthecells.

Geometry

• Coverwellswiththesuppliedlid.   Incubateat37°C

and5%CO2asusual.

12wellChamber,removableprovidesastandardslidefor-

mataccordingtoISO8037/1.

Undemanding cells can be left in their seeding medium

foruptothreedaysandgrowtoconﬂuencethere.     How-

ever, best results might be achieved when the medium

is changed every 1–2 days. Carefully aspirate the old

mediumandreplaceitby250µl/wellfreshmedium.

Geometryofthe12wellChamber,removable

Numberofwells

12

Dimensionofwells(w×l×h)    7.5×7.5×8

inmm

Volumeperwell

250µl

Totalheightwithlid

Growthareaperwell

Coatingareaperwell

Bottommaterial,sizeinmm

11mm

0.56cm2

1.9cm2

standard glass  slide

26 × 76 × 1, ground

edges,

twin  frosted

ends13mm

SurfacesandCoatings

12wellChamber,removableismountedonanuncoated

glass slide with twin frosted ends. Speciﬁc coatings on

glassarepossiblefollowingthisprotocol:

• Prepareyourcoatingsolutionaccordingtotheman-

ufacturer’sspeciﬁcationsorreference.

SolventsforFixation,StainingandOtherPur-

poses

• Apply250µlperwellandleaveatroomtemperature

foratleast30minutes.

12 well Chamber, removable is compatible to methanol,

acetone,acids,alkalis,PFA,DMSO,siliconeoil,andmin-

eraloilforcellculture.

• Aspiratethesolutionandwashwithultra-purewa-

ter.Letdryatroomtemperature.

12wellChamber,removable

Page1

Version2014-12-10

Instructions

12wellChamber,removable

ImmunoﬂuorescenceMicroscopy

• Mountslideandcoverslipwithapermanentmount-

ingmediumofyourchoiceandacoverslip24mm×

60mm(10811)(3).

Aftercultivation,cellscanbeﬁxedandstainedbeforeor

afterremovingthesiliconegasket.

Singlewelltechnique:

• Carryoutallnecessarysteps(ﬁxation,permeabiliza-

tion,staining,washing)insinglewells(1).

• Starting from one edge, remove the silicone gasket

byhandortweezers(2).

• Mountslidewithapermanentmountingmediumof

yourchoiceanda24mm×60mmcoverslip(10811)

(3).

(4)

ibidiMountingMediumisnotrecommendedfor12well

Chamber, removable because it is non–hardening and

staysaliquid(whichisadvantageousforµ–Slidesandµ–

Dishes).

Tips:

(1)

Thedaybeforeseedingthecellswerecommendplac-

ingthecellmediumandslidesintotheincubatorfor

equilibration. Thiswillpreventtheliquidinsidethe

slideorchannelfromcreatingairbubblesoverthein-

cubationtime.

Formountingofslidesamples,apermanentmount-

ing medium is  recommended.

ibidi Mounting

Medium is not recommended because it is non–

hardeningandstaysaliquid(whichisadvantageous

forµ–Slidesandµ–Dishes).

(2)

ShippingandStorage

The µ–Slides, µ–Dishes and µ–Plates are sterilized and

weldedinagas-permeablepackaging.Theshelflifeunder

proper storage conditions (in a dry place, no direct sun-

light)islistedinthefollowingtable.

Conditions

(3)

Shippingconditions

Storageconditions

Ambient

Paralleltechnique:

RT(15-25°C)

• Starting from one edge, remove the silicone gasket

byhandortweezers(2).

ShelfLifeofDifferentSurfaces

• Carryoutallnecessarysteps(ﬁxation,permeabiliza-

tion,staining,washing)dippingtheslideintotheso-

lutions. Acoverslip(24mm×60mm)canbeused

duringstainingforreducingthevolume(4).

ibiTreat,glassbottom,ESS

36months

18months

4months

Collagen,Poly-Lysine

Fibronectin

12wellChamber,removable

Page2

Version2014-12-10

Instructions

12wellChamber,removable

12wellChamber,removableFamily

Cat.No.

Description

Characteristics

81201

10811

12wellChamber,removable:microscopyglassslide

sterilized

unsterile

Coverslipsfor12wellChamber,removable,#1.5H(170± 5µm)

D263MSchottglass,24mm×60mm

12wellChamber,removable

Page3

Version2014-12-10

Instructions

12wellChamber,removable

Forresearchuseonly!

FurthertechnicalspeciﬁcationscanbefoundatForquestionsandsuggestionspleasecontactusbye-mailinfo@ibidi.de

orbytelephone+49(0)89/52046170.AllproductsaredevelopedandproducedinGermany.

©ibidiGmbH,AmKlopferspitz19,82152Martinsried,Germany.

12wellChamber,removable

Page4

Version2014-12-10